## Slide 1
Overview of
ITER Hard X-Ray Monitor
Santosh P. Pandya1,2, Robin Barnsley2, Richard O'Connor2,
Govindarajan Jagannathan2, Roger Huxford3 and Michael Walsh2
Disclaimer: The views and opinions expressed herein do not necessarily reflect those of the ITER Organization
1Aix-Marseille Université, Marseille, France
2ITER Organization, Route de Vinon-sur-Verdon, CS 90 046, 13067, St. Paul Lez Durance Cedex, France
3RBH Optics, 26 Park Close, Burgess Hill, West Sussex, RH15 8HL, UK
5th Runaway Electron Meeting (REM-2017)

## Slide 2
Hard X-Ray Monitor for ITER
Hard X-Ray Monitor is proposed for the first plasma phase in ITER (2007)
The system will detect existence of runaway electrons during plasma discharge by detecting thin-target and thick-target Bremsstrahlung emission
Bremsstrahlung
emission
Runaway electron
distribution

## Slide 3
MCNP analysis
RE current striking point
HXRM
Detector location
RE current striking point
Ref.: IDM_UID_NEFNEU
Credit: Rajnikant Makwana
Notes: 3

## Slide 4
MCNP analysis
Chart: Normalized angular flux distribution profile of the HXR photons (energy integrated)
| -180 | -160 | -140 | -120 | -100 | -80 | -60 | -40 | -20 | 0 | 20 | 40 | 60 | 80 | 100 | 120 | 140 | 160 | 180
 | 1.4185e-05 | 1.0853e-05 | 3.456e-05 | 0.00010062 | 0.00044585 | 0.00059761 | 0.00078166 | 0.0013203 | 0.0017409 | 0.98965 | 0.0018389 | 0.0014247 | 0.0007822 | 0.00060007 | 0.00047793 | 0.00011883 | 2.2066e-05 | 2.5559e-05 | 1.4185e-05
 | 1.4185e-05 | 1.0853e-05 | 3.456e-05 | 0.00010062 | 0.00044585 | 0.00059761 | 0.00078166 | 0.0013203 | 0.0017409 | 0.98965 | 0.0018389 | 0.0014247 | 0.0007822 | 0.00060007 | 0.00047793 | 0.00011883 | 2.2066e-05 | 2.5559e-05 | 1.4185e-05
 | 1.4185e-05 | 1.0853e-05 | 3.456e-05 | 0.00010062 | 0.00044585 | 0.00059761 | 0.00078166 | 0.0013203 | 0.0017409 | 0.98965 | 0.0018389 | 0.0014247 | 0.0007822 | 0.00060007 | 0.00047793 | 0.00011883 | 2.2066e-05 | 2.5559e-05 | 1.4185e-05
 | 1.4185e-05 | 1.0853e-05 | 3.456e-05 | 0.00010062 | 0.00044585 | 0.00059761 | 0.00078166 | 0.0013203 | 0.0017409 | 0.98965 | 0.0018389 | 0.0014247 | 0.0007822 | 0.00060007 | 0.00047793 | 0.00011883 | 2.2066e-05 | 2.5559e-05 | 1.4185e-05
Normalized Flux where RE beam is striking
Normalized Flux at HXRM Detector location
Notes: 4

## Slide 5
High attenuation of HXRs from FW shielding
HXR-monitor: Relatively a simple diagnostic in other tokamaks
Unlike other tokamaks, ITER has heavy shielding material and attenuates HXR-flux
Port Plug
HXR-Data
In case of other tokamaks

## Slide 6
High attenuation of HXRs from FW shielding
Penetration thickness outside Vacuum Vessel (VV): ~600 mm
Diagnostic Port Plug : ~3000 mm
DFW:~255 mm
Port Plug
Location: outside the VV

## Slide 7
Location of the Hard X-Ray Monitor
ECRH-mirror
DRS
TL5
TL3
TL9
TL7
Sector 6
Location of
HXR-monitor
(EQP12)
IRE
RE current direction
Anticipated locations
of runaway interaction

## Slide 8
High attenuation of HXRs from FW shielding
Port Plug
Potential location for HXR-monitor:
Design complications :
High temperature environment
Vacuum environment
Interface problems
Nuclear heating
PMT needs to be de-coupled from the Scintillator
Introduction of optics
Safety important penetrations

## Slide 9
Hard X-Ray system layout
Scintillator
Standard
Optical window
Lens
assembly
Optical fiber
Bundle in ISS
Bio-shield
dogleg
Optical fiber
Bundle in PCSS
Electrical Cables to
interface with DAQ
Optical fiber
Bundle inside
the DSM
PMT-box

## Slide 10
Potential Scintillator-crystals for the Hard X-Ray monitor
Properties | Scintillators |  |  |  |  |  |  | 
 | LaBr3(Ce) | CeBr3 | La-GPS | Ce:GAGG | HR-GAGG | Ce:LYSO | BGO | SrI2(Eu)
Light yield (Photons/MeV) | 63000 | 60000 | 30000 – 40000 | 56000 - 57000 | 40000 - 50000 | 34000 | 8000 | 70000 – 80000
Decay time (ns) | 16 | 19 | 50 – 70 | 88(91%) 258(9%) | 138ns(71%) 649ns(29%) | 40 | 300 | 300 – 1500
Density (g/cm3) | 5.08 | 5.10 | 5.3 | 6.63 | 6.3 | 7.1 | 7.13 | 4.6
Peak emission (nm) | 380 | 380 | 390-410 | 520 | 520 | 420 | 480 | 420
Energy resolution (% @662keV) | 2.9 | 3.8 | 5-6 | 5.2 | ~4 | 10 | 12 | 3 – 4
Hygroscopic | Yes | Yes | No | No | No | No | No | Yes
Melting point (⁰C) | 783 | 722 | 1750 | 1850 | 1850 | 2150 | - | -
Boiling point (⁰C) | 1577 | 1457 | - | - | - | - | - | -
La-GPS=Ce:(La,Gd)2Si2O7
Ce:GAGG=Ce:Gd3Al2Ga3O12
HR-GAGG=Ce:Gd3(GaAl)5O12
BGO=Bi4Ge3O12
Potential scintillator (1nd option):
Operational temperature: up to 175⁰C
Experimentally tested up to 200⁰C with heating / cooling rate 2.25 ⁰C/min
2nd Option

## Slide 11
Light outputs from Scintillator (LaBr3 and La-GPS)
Operational
regime

## Slide 12
Optical design analysis
Non-imaging collection optics having overall transmission efficiency >0.5%
Scintillator
Optical
Fiber
bundle
Collection and
transmission efficiency~95%
Optical
Fiber
bundle
Connector
Port-cell area

## Slide 13
Lowest detectable energy (in MeV) from the instrument
LaBr3(Ce)
≥60 keV, Energy detection
threshold for single
HXR-photon
Noise level
Single
HXR-photon
Direct
Coupling to PMT
Optical fiber
Coupling to PMT

## Slide 14
Credit: Julio for helping in FEM simulation
Detector size optimization for the nuclear heating
Credit: Raquel García, Juan-Pablo Catalan and Luciano Bertalot for Neutronic analysis
Chart: Steady state temperature
| 100 | 150 | 200 | 250 | 300
Scintillator (75x75) | 844 | 793 | 733 | 665 | 610
Optical Fiber Window (75x75) | 381 | 364 | 344 | 321 | 303
Optical Fiber Bundle (75x75) | 227 | 221 | 214 | 207 | 201
Scintillator (40x75) | 370 | 356 | 338 | 319 | 303
Optical Fiber Window (40x75) | 199 | 195 | 191 | 187 | 184
Optical Fiber Bundle (40x75) | 248 | 241 | 233 | 225 | 218

## Slide 15
Model runaway distribution function an MCNP calculation
Emax = 20 MeV

## Slide 16
Model runaway distribution function an MCNP calculation
Survived Flux at DFW=30 cm
Attenuation
through DFW

## Slide 17
Possible dynamic range for the measurements
At DFW=30 cm
Saturation limit
in counting mode
PMT-in current
mode
PMT-in Counting
mode
Suggested 10 kA
Detection threshold

## Slide 18
Measurement specification from HXR-monitor
SR. No. | System specifications | Value / compliance
1 | Mode of the operation | Counting mode Current mode
2 | HXR Energy range to be detected (EHXR) | ~0.5 MeV to 20 MeV
3 | HXR counts time resolution (Δt) | 1 ms to 10 ms
4 | HXR Energy spectra (dN/dE) Pulse Height Analysis of recorded counts | ~0.5 MeV to 20 MeV
5 | HXR Energy resolution (ΔE/E) | ~40% (Energy Range: 0.5 ≤ EHXR<1 MeV) ~20% (Energy Range: 1 ≤ EHXR≤20 MeV)
6 | HXR Energy spectra time resolution Δt for dN/dE | ≥10 ms (depending upon counting statistics)
7 | Maximum count rate (in counting-mode configuration) | 1.0e+6 CPS Up to 1.0e+7 CPS (post analysis algorithm)
8 | HXR-signal (in current-mode configuration) | Voltage signal showing existence of runaway electrons in plasma

## Slide 19
Summary
Preliminary design analysis of the ITER HXR-Monitor is carried out which is a first plasma diagnostic
The proposed design meets the measurement requirements within the reasonable margins and has a broad dynamic range
Further detailed design and optimization work is under progress for subsequent design phase

## Slide 20
Thank you